Notes on the calibration of the thermisters used in the CMS ECAL end cap.

Procedure

The thermisters are mounted in a cooled refrigerator on a thermally massive aluminium block.

The thermisters on the block are shielded from drafts with a polystyrene cover.

The temperature of the block is monitored using a Pt1000 resistor which is monitored by a digital voltmeter.

The refrigerator is cooled to about 5°C. It is then is turned off and allowed to warm slowly to a bit over 20C (the room is kept at 23C to 25C and heat leakage from the room provides adequate heat).

.

Data taking

Data sets are taken regularly throughout the warm up.

To take a data set takes about 60seconds and there is a 10 second pause between sets.

A data set consists of

5 consecutive readings of the Pt1000

5 consecutive readings of the first thermister

5 consecutive readings of the Pt1000

5 consecutive readings of the second thermister

5 consecutive readings of the Pt1000

5 consecutive readings of the last thermister

This process takes about 60 sec for 8 thermisters

There is then a pause, set to 10 sec

The process then repeats.

Data form

The data is presented in a text file

The name of the file has the form  yymmdd-hhmmssThermalDataFile where yymmdd is the date data taking started and hhmmss the time.

The time is in seconds from the start of the run.

The channel numbers correspond to the Pt1000 and thermister channels.

The pt1000 channel is 101

The thermister channels are 102, 103, … nnn.

The specific thermisters associated with specific channels and runs are listed in separate data files

Typically runs have been with 9 to 10 thermisters.

The thermister in channel 103 has so far been the same one to ensure cross calibration is possible in the eventuality of a Pt1000 failure

For 3 thermisters, the data would have the form:

| A | B | C | D | E | F | G |
| --- | --- | --- | --- | --- | --- | --- |
| Time | Channel | Reading 1 | Reading 1 | Reading 1 | Reading 1 | Reading 5 |
|  | 101 |  |  |  |  |  |
|  | 102 |  |  |  |  |  |
|  | 101 |  |  |  |  |  |
|  | 103 |  |  |  |  |  |
|  | 101 |  |  |  |  |  |
|  | 104 |  |  |  |  |  |
|  | 101 |  |  |  |  |  |
|  | 102 |  |  |  |  |  |
|  | 101 |  |  |  |  |  |
|  | 103 |  |  |  |  |  |
|  | 101 |  |  |  |  |  |
|  | 104 |  |  |  |  |  |

2 cycles of readings are shown.

Analysis

- The 5 readings are averaged  and the data put in column H

- The real time is entered into column A using the data from the file name

- The data is re-structured in the form

| A | B | C | D | E | F | G |
| --- | --- | --- | --- | --- | --- | --- |
| Time | Pt1000-Th102 | Th102 resistance | Pt1000-Th103 | Th103 resistance | Pt1000-Th104 | Th104 resistance |

Where the time of the start of each data set is associated with the set readings and Pt1000-Th102 is the Pt1000 reading just before the resistance of Th102 was noted, Th102 resistance being the associated resistance.

The temperature of the Pt1000 was then calculated using the equation.

T=−A+A2+4B1−RtR02B

Which is derived from the equation

Rt=R01+At+Bt2

Where 	A= 3.9083E-3

B= -5.775E-7

Rt= the resistance at temperature T

R0 = the resistance at 0°C

The temperature given by this parameterisation are in °C

The Pt1000 temperatures were then associated with the relevant thermisters.

The Steinhart-hart equation was then used to parameterise the variation of resistance with temperature of the thermisters.

The Steinhart-hart equation has the form

1t=A+BLnR+CLnR3

Values of A, B and C were determined for each thermister by using observed values for resistance and temperature at three temperatures, typically, as near as possible to 10°C, 15°C and 20°C.

This was then used to calculate the predicted temperature over the whole temperature range.

The biggest error observed was 0.017°C