Notes on calibration of thermisters.

Calibration runs have been run for all the “PR” series thermisters and the “t” series  thermisters up to t_374

The data on the PR series that have been installed in the test “D” has been fitted to a binomial parameterisation

This parameterisation is provisional as it has been found that there is a small time lag between the measured temperature and the actual temperature of the thermisters during tests and this needs to be corrected for.

There was a problem during an early run with the “t” series thermisters and some of the calibrations failed.

Provisional fits for thermisters installed on the test D

In the following table, x = resistance, y temperature.

| Sasha'a list of thermisters in the test beam |  |  |
| --- | --- | --- |
| Thermister | Callibration Run | Binomial Fit, 15 to 20 C |
| PR 001 | 060608 | y = 6.5005E-10x2 - 3.4136E-04x + 5.2477E+01 |
| PR 002 | 060608 | y = 6.3064E-10x2 - 3.3573E-04x + 5.2288E+01 |
| PR 003 | 060608 | y = 6.5725E-10x2 - 3.4251E-04x + 5.2343E+01 |
| PR 004 | 060608 | y = 6.3066E-10x2 - 3.3594E-04x + 5.2457E+01 |
| PR 005 | 060608 | y = 6.6964E-10x2 - 3.4569E-04x + 5.2392E+01 |
| PR 006 | 060608 | y = 6.5871E-10x2 - 3.4325E-04x + 5.2375E+01 |
| PR 007 | 060608 | y = 6.2296E-10x2 - 3.3262E-04x + 5.2069E+01 |
| PR 009 | 060608 | y = 6.3251E-10x2 - 3.3587E-04x + 5.2265E+01 |
| PR 012 | 060705 | y = 6.5829E-10x2 - 3.4423E-04x + 5.2754E+01 |
| PR 013 | 060705 | y = 6.5473E-10x2 - 3.4294E-04x + 5.2719E+01 |
| PR 014 | 060705 | y = 6.5400E-10x2 - 3.4155E-04x + 5.2365E+01 |
| PR 015 | 060705 | y = 6.8458E-10x2 - 3.5152E-04x + 5.2929E+01 |
| PR 020 | 060914 | y = 6.6580E-10x2 - 3.4375E-04x + 5.2038E+01 |
| PR 021 | 060914 | y = 6.4525E-10x2 - 3.3890E-04x + 5.2203E+01 |
| PR 024 | 060914 | y = 6.2924E-10x2 - 3.3477E-04x + 5.2240E+01 |
| PR 026 | 060914 | y = 6.4202E-10x2 - 3.3800E-04x + 5.2224E+01 |
| PR 029 | 060918 | y = 6.7697E-10x2 - 3.4878E-04x + 5.2775E+01 |
| PR 031 | 060918 | y = 6.7808E-10x2 - 3.4954E-04x + 5.2909E+01 |
| PR 032 | 060918 | y = 6.9964E-10x2 - 3.5622E-04x + 5.3297E+01 |
| PR 035 | 060921 | y = 7.1369E-10x2 - 3.5902E-04x + 5.3063E+01 |

Choice of Parameterisation

A binomial parameterisation was chosen as it is simpler to apply than the more complex Steinhart-Hart Equation and gave a good fit over the limited temperature range from 15°C to 20°C

Fig.1 shows a comparison of a fit between 10 and 20°C, compared to a fit between 15 and 20°C. It was found necessary to express the constants to 5 figures.

Variations between thermisters

Fig 2 shows the variation in resistance between thermisters for 7 different thermisters measured at the same time.

Problems associated with rate of rise of temperature in test assembly

Fig 3 shows small variations in the measured characteristics of the same thermister t_023 for different calibration runs

A plot of resistance against heating rate Fig. 4 shows a correlation

With further analysis it should be possible to correct for the heating rate error.

Data

There is good data on thermisters PR01 to PR042

A problem was found in the testing rig when it was first brought out to CERN

This has been rectified but in the data for t_121 to t_205 there are some problems with a few measurements (t_017 and t_019 are stray thermisters for which we have good data).

There is no data on the following 15 thermisters,

t_129, t_130, t_157, t_190 to t_200 and t_205

There is suspect data on the following 2 thermisters,

t_173 and t_156

There is compromised data, possibly recoverable on the following 16 thermisters t_141 to t_155 and, t_180

All other data should be good, as is any data for t_206 and above.

So far we have measured up to t_374.

Data Storage

I sent copies of the early data to you at CERN.

The later data is backed up by Julian on the Flint drive

Data Analysis

I am using an excel program to analyse the data

P.S.Flower

11 Sept 07